Lot Acceptance Report – L23a

STMicroelectronics

Laser Transmitters

Internal Report

R. Zuyeuski

Table of contents

1.	Lot Validation Program	3

2.	Technical validation	4

2.1	Compliance with technical requirements	4

2.2	Validation test results	4

2.2.1	Visual Inspection	4

2.2.2	LIV Characteristics	5

2.2.3	Dynamic Measurement	7

3.	Conclusions	8

# Lot Validation Program

This is an internal lot validation report for the Laser Transmitters (Laserpills) supplied by STMicroelectronics under the terms of CERN Contract F 451/EP/CMS. The lot validation testing flow is shown in Figure 1. The figure shows the tests to be carried out as described in the test procedure for laserpills (Laser Transmitter Quality Assurance Procedure; CMS-TK-QP-0002). The test target specifications are extracted from the technical specification for delivered laserpills (CMS-TK-ES-0004).

Figure 1.	Lot validation testing flow

40 lasers, or 5% of total number, of lot L23a delivered by ST in June 2004 were sampled randomly and tested (see Table 1). All tests were carried out at room temperature.

Table 1.	ST Serial Numbers of the 40 devices tested at CERN and their correpsonding CMS Serial Numbers

| ST Serial # | CMS Serial # | ST Serial # | CMS Serial # |
| --- | --- | --- | --- |
| XI446B40 | 30201010043882 | XI455B17 | 30201010038662 |
| XI446B63 | 30201010043491 | XI455B25 | 30201010038355 |
| XI447B35 | 30201010043885 | XI455B38 | 30201010039095 |
| XI450B26 | 30201010039824 | XI455B76 | 30201010039472 |
| XI450B60 | 30201010038408 | XI459B01 | 30201010043748 |
| XI450B66 | 30201010043275 | XI459B25 | 30201010039493 |
| XI450B79 | 30201010038463 | XI459B26 | 30201010039446 |
| XI450B92 | 30201010038721 | XI459B27 | 30201010039828 |
| XI452B04 | 30201010039236 | XI459B28 | 30201010040244 |
| XI452B14 | 30201010039458 | XI459B30 | 30201010040293 |
| XI452B17 | 30201010040222 | XI459B34 | 30201010039284 |
| XI452B27 | 30201010040386 | XI459B47 | 30201010038571 |
| XI454B09 | 30201010038351 | XI460B16 | 30201010040280 |
| XI454B13 | 30201010038377 | XI461B05 | 30201010038319 |
| XI454B19 | 30201010043455 | XI461B40 | 30201010038752 |
| XI454B27 | 30201010038591 | XI461B42 | 30201010038257 |
| XI454B62 | 30201010040343 | XI461B50 | 30201010015576 |
| XI455B04 | 30201010043873 | XI461B61 | 30201010043173 |
| XI455B14 | 30201010043359 | XI461B75 | 30201010043437 |
| XI455B16 | 30201010043624 | XI461B90 | 30201010043364 |

# Technical validation

## Compliance with technical requirements

Table 2 shows the specified performance of the laser transmitters against which the received sample was tested. The table also gives a summary of the results obtained in comparison with the specifications.

Table 2.	Specification test targets taken from the technical specification

| Specification to be tested | Test Target Specifications | CERN Validation Results |
| --- | --- | --- |
|  | Min | Max | Units | Min | Max |
| Threshold current |  | 10 | mA | 5.880 | 6.603 |
| Slope efficiency | 0.032 | 0.048 | mW/mA | 0.036 | 0.047 |
| Forward voltage |  | 1.75 | V | 1.503 | 1.570 |
| Fibre output power | 1.3 | 1.9 | mW | 1.409 | 1.873 |
| Integral linearity deviation |  | 1 | % | 0.071 | 0.517 |
| Rise time |  | 0.5 | ns | 0.226 | 0.383 |

## Validation test results

### Visual Inspection

All 800 devices from lot L23a have polyethylene-buffered fibre. 40 inspected lasers have no visible defects except for 4 cases (ST IDs XI459B27, XI459B28, XI454B19 and XI461B75), where relatively big spots of glue were observed at the bottom side (see Figure 2) that could create assembling problem. 40 more devices were inspected for the glue, and in 3 cases (ST IDs XI466B00, XI447B89 and ID XI458B56) the same problem was observed. In all cases glue spots are close to the centre of laser. One of such lasers (ST ID XI458B56) was given to ST Microelectronics for investigation.

Figure 2.	Lasers contaminated with glue (a) and epoxy (b).

One device (ST ID XI457B03) has silver epoxy or paint on the buffer (see Figure 2b), ST promised to investigate it also.

In total, 8 devices out of 80 did not pass visual inspection.

### LIV Characteristics

Figure 3, Figure 4 and Figure 5 show the results of the LIV measurements performed at CERN. The colored lines indicate the range or maximum value allowed for each parameter. From the plots is clear that all lasers met the criteria given in the specifications.

Figure 3.	Efficiency and threshold for the 40 devices measured at CERN

Figure 4.	Forward voltage and fiber output power for the 40 devices measured at CERN

Figure 5.	Maximum value of the integral linearity deviation for the 40 devices measured at CERN

Figure 6 and Figure 7 show comparisons between CERN and ST measurements of Ith, Slope Efficiency and Forward Voltage. The scatter plots in Figure 6 show that there is a good correlation between ST and CERN measurements even if the measurement of the threshold performed at CERN is always bigger than the corresponding ST value. This could be due to the fact that the power was a bit underestimated by the CERN set up. Nevertheless, since the difference is not big (5.7% in average), and the CERN measurements are well inside the specifications we can accept the lasers.

Figure 6.	Comparison of CERN-measured Ith and efficiency with the ST measurements.

Figure 7.	Comparison of CERN-measured forward voltage with the ST measurements.

The spread in the distribution on Figure 7 shows that the measurements of the forward voltage are stable, but always bigger than the corresponding ST values. And again, since the difference is not big (4.2% in average), and the CERN measurements are well inside the specifications we can accept the lasers.

### Dynamic Measurement

By applying a voltage pulse train to the laser under test it is possible to excite the laser to laser from below threshold. By observing the resultant optical pulse on an oscilloscope with the aid of an optical probe one can measure the relaxation oscillations of the laser die and thus determine the intrinsic bandwidth of the die without the sub-mount. A typical optical pulse is shown in Figure 8, while the extracted bandwidth and risetime are shown in Figure 9

Figure 8.	Optical output pulse as measured on an oscilloscope with optical probe. The laser bandwidth is determined from the relaxation oscillation frequency.

Figure 9.	Bandwidth and risetime for the 40 devices that were measured at CERN. All devices are in spec: Bandwidth>1GHz.

# Conclusions

All 40 lasers tested at CERN were found inside specification. At the same time, 8 devices out of 80 did not pass visual inspection. Therefore we decided to accept the lot but hold it till results of ST investigation.

# Footnotes

Laser Transmitter Quality Assurance Procedure, Version 1.2, 17 January 2003

CMS Tracker Optical Readout Link Specification, Part 2.2: Laser Transmitter, Version 3.5, 14 May 2002

# Text boxes

b

a

epoxy

glue