MPO Terminated long ribbon

Lot acceptance Report

15 February 2005

# Overview

In this document we report on the lot acceptance tests carried out on 13 randomly sampled long ribbons terminated with MPO connectors. The samples were delivered as batch . Table 1 shows the sample history and Table 2 the list of the samples.

| Batch delivery | Sumitomo Harnesses Batch | Delivery Date | Quantity Delivered | Length (m) | Quantity Sampled | Diamond QR-MPO-code 33013064200XXX |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 77 76 50 115 110 83 20 74 - 46 | 18 19 20 21 23 25 26 27 28 40 | 2 2 2 1 1 1 1 1 1 1 | 864 - 083 220 - 058 042 - 282 183 123 238 399 388 303 340 |

Table 1: sample history

| Diamond QR-MPO-code 33013064200XXX | Diamond QR-MPO code 33013064100XXX |
| --- | --- |
| 042 | 059 |
| 123 | 845 |
| 183 | 186 |
| 058 | 085 |
| 083 | 005 |
| 388 | 489 |
| 399 | 250 |
| 303 | 463 |
| 282 | 323 |
| 220 | 305 |
| 864 | 119 |
| 238 | 111 |
| 340 | 438 |

Table 2: List of MPO terminated samples

# Production tests and results

The procedure for production qualification is described in document “Short fanout and long ribbon Technical specifications Version 1.00”. The sequence followed for this acceptance test is illustrated in Figure 1. The specifications tested are shown in Table 3.

The 13 selected terminated fibre ribbons underwent visual inspection, geometrical tests and optical tests.

.

Figure 1:  Flow chart of the procedure.

| # | Specification to be tested | Test Specifications | CERN Testing |
| --- | --- | --- | --- |
|  |  | Min | Typ | Max | Units | Production Acceptance |
| 3.1 | Number of channels | 12 |  | ( |
| 3.2 | Operation Wavelength | 1310 | nm |  |
| 3.3 | Connector type (multi-fibre) | MPO |  | ( |
| 3.4 | Number of mating cycles | 20 |  |  |  |  |
| 3.5 | Random mate insertion loss |  |  | 1.2 | dB | ( |
| 3.6 | Return loss | 45 |  |  | dB | ( |
| 3.7 | Tensile load on connector |  |  | 10 | N |  |
| 3.8 | Ribbon length Lribbon | Nominal length -0/+10 | cm | ( |

Legend:

( Visual Inspection	( Geometrical measurement	( Optical Test

Table 3:  Validation Programme overview.

## Visual inspection

- All samples were terminated correctly with MPO connectors as specified.

- The packaging by Diamond is acceptable and the coil diameter is in specification.

- The QR labels were properly attached according to the specification and could be read without trouble.

Nevertheless one label has been removed during manipulation: the labels on connector body should be glue stronger.

- The ribbon orientation into the MPO connector was as specified.

## Geometrical measurement

All fibres of each sample were measured. The total length was measured with a high resolution OTDR for one channel. Table 4 shows that all ribbons are within specifications for the nominal length.

| Long ribbon QR code | Length (m) |
| --- | --- |
| 042-059 | 20.38 |
| 123-845 | 23.32 |
| 183-186 | 21.39 |
| 058-085 | 19.28 |
| 083-005 | 18.37 |
| 388-489 | 27.37 |
| 399-250 | 26.40 |
| 303-463 | 28.37 |
| 282-323 | 20.21 |
| 220-305 | 19.36 |
| 864-119 | 18.39 |
| 238-111 | 25.34 |
| 340-438 | 40.34 |

Table 4: length of MPO terminated ribbon

## Optical measurements

Optical properties were measured using the setup shown in Fig. 2. We note that each MU connector insertion loss was measured. Thus there are 2 optical measurements values for one ribbon.

Fig.2: Schematic measurement setup for optical measurements.

2.3.1.	Random mate insertion loss (all fibre channels).

The insertion loss (IL) is defined as the Log of the ratio of optical powers measured before (P0) and after (P1) insertion of the device under test.  The optical powers (P0 and P1) were measured with a large area detector in an optical head receptacle.

Insertion loss:

#### 2.3.1.a	Acceptance criteria

The acceptance criterion is such that the insertion loss of the MFS connector shall be less than 1.2 dB.

#### 2.3.1.b	Results

The results are shown in Fig. 3: all the fanout ribbons were found to be within the specified limit.

Fig.3: Insertion loss of MPO terminated ribbons

### 2.3.2.	Return loss (all fibre channels).

The return loss (RL) is defined as the Log of the ratio of the reflected optical powers measured during referencing (Pref) and after (P1) insertion of the device under test (DUT).

Return loss: (1)

#### 2.3.2.a	Acceptance criteria

The acceptance criterion is such that the return loss shall be greater than 50dB.

#### 2.3.2.b	Results

The results are shown in Fig. 4. All devices were well within the specified limit.

Fig.4: Return loss of MPO terminated ribbons.

# Conclusions

The 13 MPO terminated long ribbons sampled are in specification.

All  ribbons supplied by Diamond as part of batch XXX are thus accepted.

Note that QR code labels are not so well attached with this type of glue.

# Footnotes

(1) Formula valid if using a reference with perfect glass-air interface.