# MPO Multi-Ribbon Cable

# Technical Specifications

# Version 2.4

Technical contacts:

Alexander Singovski, Univ. Minnesota

CERN, Div. PH/UCM

Alexander.Singovski@cern.ch

+41 (022) 767 1674

In case of absence of A.Singovski:

François Vasey, CERN

CERN, Div. PH/CME

François.Vasey@Cern.ch

+41 (022) 767 38 85

and

Wolfgang Funk, CERN

CERN, Div. PH/CMA

Wolfgang.Funk@Cern.ch

+41 (022) 767 33 76

Table of Contents

1 .  Background	3

2 .  MPO Multi-Ribbon Cable	4

3 .  Dust Covers for Connectors and Adaptors	5

4 .  Technical Alternative for the In-line adaptor	5

5 .  Materials Required and Free-Issue Materials	5

6 .  Suggested Manufacturing Sequence	6

7 .  MPO Multi-Ribbon Cable Specifications	7

7.1 Operational Specifications	7

7.2 Geometric Specifications	7

7.3 Environmental Specifications	8

7.4 Safety Specifications	8

7.5 Other technical Characteristics	8

7.5.1 Test documentation and traceability	8

7.5.2 Shipping and storage requirements	9

7.6 Testing	9

# .  Background

The University of Minnesota (Minnesota, U.S.) and CERN (European Organization for Nuclear Research) intend to procure 8-ribbon, 12-fiber optical multi-ribbon cables terminated in MPO connectors and MPO optical adaptors (“in-line adaptors”) for use in the on-going CMS ECAL project at CERN, for which the University of Minnesota and CERN have the responsibility of providing the 0.8 Gb/s digital optical data link system.

The MPO multi-ribbon cable and in-line adaptor form part of the optical link fiber chain for two ECAL sub-projects: ECAL Endcap and ECAL Preshower.  (The third sub-project ECAL Barrel employs MFS-MPO multi-ribbon cables and MFS in-line adaptors [] in place of these two items.)

The place of the multi-ribbon cable and in-line adaptor within the optical link system is shown in Figure 1.

Figure 1.  Block diagram of the CMS ECAL optical data link system.

A total of approximately 106 MPO multi-ribbon cables and 260 4-way MPO in-line adaptors will be procured for ECAL.  They shall be required to operate within specifications throughout their lifetime of approximately 10 years, subjected to a harsh environment of radiation and high magnetic field.

# .  MPO Multi-Ribbon Cable

The following description and comments from Reference [1] apply to the MPO multi-ribbon cable as well:

“The dense multi-ribbon cable consists of a bundle of eight 12-way ribbons protected by a low smoke, zero halogen sheath reinforced by two lateral strength-members. At both ends, the cable breaks out into 8 naked ribbons to be terminated with MT-based connectors. On the in-line patch-panel side, connectors must be non-magnetic of [type MPO12] without guide pins. On the backend patch-panel side, connectors shall be of the MPO12 type, with guide pins. The length of the assembly will be approximately 60m. Once installed, the cable will cross the concrete shield separating the detector volume from the counting room. It will be pulled manually across a cable duct and should thus withstand standard installation and operation conditions.

The dense multi-ribbon cable is specified in part 4.1 (Reference [])…and the back-end patch-panel in part 8 of the CMS Tracker Optical Readout Link specification document.”

A diagram of the terminated MPO multi-ribbon cable is shown in Figure 2.

Figure 2.  Diagram of the terminated MPO multi-ribbon cable.

The free-issue multi-ribbon cable has length markings in meters running in a line down the outside of the sheath.  The backend part of the assembly shall be performed on the cable end having the lower meter-marking value on the sheath.

The assembly of the terminated MPO multi-ribbon cable includes:

breakout and termination of ribbons on both ends

re-placement of the Ericsson clips which prevent ribbon stack rotation

placement of dust covers on all 16 connectors

assembly of the flexible protective sleeve on the backend

fixation of one plate-type QR code label on each end of the Ericsson sheath

gluing of one die-type QR code label on each backend connector

A detailed diagram of the flexible sleeve assembly over the broken-out ribbons at the backend is shown in Figure 3.  The sleeve material is fixed with sturdy heat-shrink tubing, over the sheath of the raw Ericsson cable on the sheath end, and over a brass ferrule on the bare-ribbon end.  A short piece of thin tubing is placed around the ribbon stack to keep it from separating inside the flexible sleeve.

Figure 3.  Breakaway view of the flexible sleeve assembly at the backend of the MPO multi-ribbon cable.

# .  Dust Covers for Connectors and Adaptors

The manufacturer will furnish each MPO connector with one dust cover in place, and each adaptor with one dust cover in place.

# .  Technical Alternative for the In-line adaptor

The University of Minnesota and CERN reserve the right to require the manufacturer to substitute the 260 4-way MPO in-line adaptors and dust covers with 1,000 1-way MPO adaptors and dust covers.

# .  Materials Required and Free-Issue Materials

The following materials are required to assemble one MPO multi-ribbon cable.  Items which will be free-issued by CERN are noted as such.

one length of raw multi-ribbon cable (free-issue, in pre-cut lengths)

four Ericsson clips to prevent ribbon rotation (free-issue, already in place)

one length of flexible sleeve material (free-issue)

two pieces of sturdy heat-shrink tubing

one brass ferrule

one length of thin tubing

materials necessary to assemble eight MPO/APC/F connectors and eight MPO/APC/M connectors

sixteen dust covers for MPO connectors

two plate-type QR code labels (free-issue)

eight die-type QR code labels (free-issue)

The raw multi-ribbon cable is specified in part 4.1 (Reference [2]) of the CMS Tracker Optical Readout Link specification document.  The flexible sleeve material is AWAG (Reference []) article no. 24.45.0052:  Hahn HPP 14/16 8.5-18 mm.

# .  Suggested Manufacturing Sequence

The following manufacturing sequence is recommended.  The manufacturer remains responsible however for the suitability of the manufacturing sequence employed.

- Orient the length of raw multi-ribbon cable so that the backend will be assembled on the end having the lower meter-marking value on the sheath.

- Slide a length of flexible sleeve material and a length of heat-shrink tubing over the backend of the raw multi-ribbon cable.

- Cut a hole in the outer sheath of the multi-ribbon cable.

- Pull ribbon through the hole, being careful to avoid separation of the ribbon stack, rotation of ribbons relative to one another, or rotation of the ribbon stack within the remaining sheath.

- Remove the Ericsson clips from the end of the sheath and cut off the excess sheath.

- Replace the clips on the newly-cut end, being careful to avoid separation of the ribbon stack, rotation of ribbons relative to one another, or rotation of the ribbon stack within the remaining sheath.

- Slide the thin tubing over the ribbon stack, being careful to avoid separation of the ribbon stack, rotation of ribbons relative to one another, or rotation of the ribbon stack within the remaining sheath.

- Apply epoxy between the sheath and the sleeve and between sleeve and the heat-shrink tubing, and apply heat to shrink the tubing, thus fixing these three parts together (at the Ericsson-clips end).

- Slide the brass ferrule over the thin tubing.

- Apply epoxy between the ferrule and the sleeve and between sleeve and the heat-shrink tubing, and apply heat to shrink the tubing, thus fixing these three parts together (at the brass ferrule end).

- Repeat steps 3 through 6 to break out the in-line end of the multi-ribbon cable.

- Terminate the ribbons on both ends in connectors.

- Apply the dust covers and the QR-code labels.

# .  MPO Multi-Ribbon Cable Specifications

The following technical specifications must be satisfied by all 96 channels of the MPO multi-ribbon cable.

All technical specifications are for assemblies at 25(C unless otherwise noted.

Technical specifications in italics apply to the free-issue raw multi-ribbon cable, described in part 4.1 (Reference [2]) of the CMS Tracker Optical Readout Link specification document. The manufacturer is not required to test specifications in italics in this procurement.

## Operational Specifications

| # | Operational Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 6.1.1 | Number of channels | 8 x12 |  |  |
| 6.1.2 | Operation wavelength | 1310 | nm |  |
| 6.2.1 | Connector type on in-line patch panel end | MPO12/APC/female |  | IEC 61754-7 without guide pins |
| 6.2.2 | Connector type on backend patch panel end | MPO12/APC/male |  | IEC 61754-7 with guide pins |
| 6.2.3 | MPO12 guide pin length |  |  | 3.0 | mm | non-removable |
| 6.2.4 | Number of mating cycles | 100 |  |  |  | all MPO connectors |
| 6.2.5 | Random mate insertion loss1 |  |  | 1.2 | dB | all channels, measured at both ends |
| 6.2.6 | Random mate return loss1 | 50 |  |  | dB | all channels, measured at both ends |
| 6.3.1 | Tensile load on connectors |  |  | 10 | N | operation and installation |
| 6.3.2 | Torsion on connectors |  |  | 5 | turns/m | installation |
|  |  |  |  | 1 | turn/m | operation |

1Insertion loss and return loss are defined in Reference [1].

## Geometric Specifications

| # | Geometric Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 6.4.1 | Sheath section length Lsheath | 50 | 60 | 70 | m | Lsheath to be distributed in a set of no more than 10 standard lengths, with tolerances +30cm/-30cm. |
| 6.4.2 | Multi-ribbon cable orientation |  |  |  |  | Backend shall have the lower meter-marking value on the sheath. |
| 6.4.3 | In-line ribbon section LRinline |  | 55 |  | cm | Naked ribbon, tolerance +1.5cm/-1.5cm |
| 6.4.4 | Backend ribbon section LRbackend | 65 |  | 235 | cm | Naked ribbon; 28 cables will have LRbackend = 235 cm. Remaining cables will have LRbackend distributed in a set of lengths staggered by 4.4cm: ribbon 1 = 90.6cm, ribbon 2 = 86.2cm, …, ribbon 8 = 59.8cm. The ribbon length tolerance: ribbon-to-ribbon =+0.5cm/-0.5cm; overall ribbon length= +1.5cm/-1.5cm |
| 6.5.1 | Flexible sleeve length LSleeve |  | 47 |  | cm | Unstretched sleeve; Tolerance +2.0cm/-2.0cm |
| 6.6.1 | Brass ferrule length |  | 3.0 |  | cm | Tolerance +0.2cm/-0.2cm |
| 6.6.2 | Brass ferrule inner diameter |  | 6.0 |  | mm | Tolerance +0.2mm/-0.2mm |
| 6.6.3 | Brass ferrule outer diameter |  |  |  |  | Not specified; the ferrule must have an outer diameter such that specification 6.6.4 is fulfilled. |
| 6.6.4 | Outer diameter of finished assembly of brass ferrule + flexible sleeve + heat shrink tubing |  | 12.0 |  | mm | Tolerance +0.5mm/-0.5mm |
| 6.6.5 | Thin tubing length |  | 5 |  | cm | Tolerance +1cm/-1cm |
| 6.6.6 | Thin tubing diameter |  | 6.0 |  | mm | Thin tubing must slide easily on ribbon stack but not allow stack to separate. Thin tubing must also slide into brass ferrule. |

## Environmental Specifications

| # | Environmental Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 6.7.1 | Magnetic field |  |  | 4 | T | Parallel to any axis |
| 6.7.2 | Hadronic fluence1 |  |  | 1x1014 | cm-2 | Integrated over lifetime2 |
| 6.7.3 | Gamma radiation dose1 |  |  | 1.5x105 | Gy (Si) | Integrated over lifetime2 |
| 6.7.4 | Temperature | 18 | 25 | 70 | °C | Operation and storage |
| 6.7.5 | Operating humidity |  | TBD |  |  |  |

1Component resistance to radiation will be verified under the sole responsibility of University of Minnesota and CERN.
2Foreseen operating lifetime:  nominal 10 years.

## Safety Specifications

| # | Safety Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 6.7.6 | Material composition | Halogen-free, flame retardant material |  | CERN IS41; see reference []. |

## Other technical Characteristics

Connector color	TBD, before contract placement.

Ribbon orientation	Index on connector side shall indicate fiber no. 1.  Numbering according to Bellcore color code (fiber 1: dark blue).

Connector boot	To be determined by University of Minnesota/CERN, before contract placement.

Dust protection	On in-line end: Adaptor shall be protected with dust cap.
On backend:  Connector shall be protected with dust cap.

Labelling	One free-issue plate-type QR-code label shall be fixed on each end of the multi-ribbon cable sheath (2 labels per assembly).
One free-issue die-type QR-code label shall be glued on each backend connector body (8 labels per assembly).

### Test documentation and traceability

Test file shall be delivered at least in electronic form.  Data for each multi-ribbon cable assembly shall include:

manufacturer serial number (if different from QR-code) (one number per assembly)

plate QR-code number corresponding to the multi-ribbon cable (one number per assembly)

multi-ribbon cable nominal sheath length Lsheath (one number per assembly)

die QR-code number corresponding to the backend MPO connector (eight numbers per assembly)

backend ribbon section nominal length LRbackend (eight numbers per assembly)

fiber number (1-12) (96 numbers per assembly)

from which end the measurement is taken (in-line or backend)

insertion loss and return loss (2x96 insertion loss and 2x96 return loss numbers per assembly)

### Shipping and storage requirements

Shipping	Packaged in free-issue Ericsson crate in which raw multi-ribbon cable length was received, with appropriate protection of naked ribbon sections.

Storage	According to best practice.

## Testing

| # | Specification to be tested | Manufacturer | Minnesota/ CERN |
| --- | --- | --- | --- |
|  |  | Product Qualification | Lot validation (before delivery) | Lot Acceptance |
| 6.1.1 | Number of channels | ♦ | ♦ | ♦ |
| 6.1.2 | Operation wavelength1 | ♦ | ♦ | ♦ |
| 6.2.1 | Connector type on in-line patch panel end | ♦ | ♦ | ♦ |
| 6.2.2 | Connector type on backend patch panel end | ♦ | ♦ | ♦ |
| 6.2.3 | MPO12 guide pin length | ♦ | ♦ | ♦ |
| 6.2.4 | Number of mating cycles | ♦ |  | ♦ |
| 6.2.5 | Random mate insertion loss2 | ♦ | ♦ | ♦ |
| 6.2.6 | Random mate return loss2 | ♦ | ♦ | ♦ |
| 6.3.1 | Tensile load on connectors | ♦ |  | ♦ |
| 6.3.2 | Torsion on connectors | ♦ |  | ♦ |
| 6.4.1 | Sheath section length Lsheath |  |  | ♦ |
| 6.4.2 | Multi-ribbon cable orientation | ♦ | ♦ | ♦ |
| 6.4.3 | In-line ribbon section LRinline | ♦ | ♦ | ♦ |
| 6.4.4 | Backend ribbon section LRbackend | ♦ | ♦ | ♦ |
| 6.5.1 | Flexible sleeve length | ♦ | ♦ | ♦ |
| 6.6.1 | Brass ferrule length | ♦ |  | ♦ |
| 6.6.2 | Brass ferrule inner diameter | ♦ |  | ♦ |
| 6.6.3 | Brass ferrule outer diameter | ♦ |  | ♦ |
| 6.6.4 | Outer diameter of finished assembly of ferrule + flexible sleeve + heat shrink tubing | ♦ | ♦ | ♦ |
| 6.6.5 | Thin tubing length | ♦ |  | ♦ |
| 6.6.6 | Thin tubing diameter | ♦ |  | ♦ |
| 6.7.1 | Magnetic field | ♦ |  | ♦ |
| 6.7.2 | Hadronic fluence |  |  | ♦ |
| 6.7.3 | Gamma radiation dose |  |  | ♦ |
| 6.7.4 | Temperature | ♦ |  |  |
| 6.7.5 | Operating humidity | ♦ |  |  |
| 6.7.6 | Material composition | ♦ | ♦ |  |

1Wavelength to be used for testing is ~1310 nm.
2Insertion loss and return loss are defined in Reference [1].

References

# Endnotes

[]  F. Vasey, “CMS Tracker Optical Readout Link Specification, Part 4: Terminated Multi-Ribbon Cable”, v. 3.5, CERN EP/CME (2002).  https://edms.cern.ch/file/351654/3.5/TerminatedMultiRibbonCableV3_5.pdf

[] F. Vasey, “CMS Tracker Optical Readout Link Specification, Part 4.1: Dense Multi-Ribbon Cable, v. 3.7, CERN EP/CME (2002).  
https://edms.cern.ch/document/312744/3.7

[]  http://www.awag.ch

[]  CERN TIS Commission, “The Use of Plastic and other Non-Metallic Materials at CERN with respect to Fire Safety and Radiation Resistance”, CERN TIS IS 41 (1995). http://edms.cern.ch/file/335806/1.01/IS41_E.pdf
http://edms.cern.ch/document/108727/2

# Text boxes

MPO12/APC/M

Dust
cover

Flexible protective sleeve over bare 12-fiber ribbon

Clips to prevent ribbon stack rotation

Bare 12-fiber ribbon

MPO12/APC/F

Plate-type 
QR code labels

Loose-tube rugged 
multi-ribbon cable

Sheath section Lsheath

12

Digital amp. ASIC

PIN photo-diode array

Off Detector

12

Rx module

Intermediate
adaptor

In-line
adaptor

1

Long
ribbon

12

Intermediate Patch Panel

MU-SR
adaptor

CMS

In-Line Patch Panel

Laser diode

Laser driver ASIC

Front End

1

Opto-hybrid Tx

Dense multi-ribbon cable

Short
fanout

Pigtail fiber

Back-end Patch Panel

Distributed Patch Panel

96

12

12

Backend bare ribbon section with distribution of values of LRbackend

...

Flexible sleeve length Lsleeve

Dust
cover

Die-type QR-code label

In-line ribbon section LRinline

...

Thin tubing to keep ribbon stack from separating

Ericsson cable sheath

Clips (placed by Ericsson)

Brass ferrule

Sturdy heat-shrink tubing

Flexible sleeve