# Short Fanout and Long Ribbon

# Technical Specifications

# Version 1.00

Document history:

Version 0.65,	 14 Jun 04		First released draft version

Version 1.00,    04 Aug 04		First released version

Technical contact:

James Grahl, Univ. Minnesota–Minneapolis

Contact information:

CERN, Div. PH

James.Grahl@Cern.ch

+41 (022) 767 83 47

Table of Contents

1 .  Background	3

2 .  Possible Construction Sequences	4

2.1 Sequence A: Construction from free-issue harnesses and ribbon	4

2.2 Sequence B: Construction from free-issue long fanouts	4

2.3 Intermediate connectors and adaptor	5

3 .  Short Fanout Specifications	7

3.1 Operational Specifications	7

3.2 Geometric Specifications	7

3.3 Environmental Specifications	8

3.4 Safety Specifications	8

3.5 Other Characteristics	8

3.5.1 Intermediate Fanout Connector	8

3.5.2 Test documentation and traceability	8

3.5.3 Shipping and storage requirements	8

3.6 Testing	9

3.7 Implementation	9

4 .  Long Ribbon Specifications	10

4.1 Operational Specifications	10

4.2 Geometric Specifications	10

4.3 Environmental Specifications	11

4.4 Safety Specifications	11

4.5 Other Characteristics	11

4.5.1 Intermediate Ribbon Connector	11

4.5.2 Test documentation and traceability	11

4.5.3 Shipping and storage requirements	11

4.6 Testing	12

4.7 Implementation	12

5 .  References	13

# .  Background

The University of Minnesota (Minnesota) intends to procure 12-fiber optical fanouts (“short fanouts”), 12-fiber optical jumper ribbons (“long ribbons”) and optical adaptors connecting them (“intermediate adaptors”) for use in the on-going CMS ECAL project at CERN, for which Minnesota and CERN have the responsibility of providing the 0.8 Gb/s digital optical data link system.

The short fanout, intermediate adaptor and long ribbon form the middle of the optical link fiber chain for ECAL sub-projects ECAL Endcap and ECAL Preshower.  (The third sub-project ECAL Barrel employs “standard fanouts” [3] in place of these three items.)  The short fanout carries light from a MU-SR adapter to the intermediate adaptor.  The long ribbon carries light from the intermediate adaptor to an MF-B2/MF-B3 adaptor (“MFS adaptor”; Diamond proprietary).

The place of the short fanout, intermediate adaptor and long ribbon within the optical link system is shown in Figure 1.

Figure 1.  Block diagram of the CMS ECAL optical data link system.

Approximately 600 short fanouts, long ribbons and intermediate adaptors will be procured for ECAL.  They will be required to operate with perfect reliability throughout their lifetime of approximately 10 years, subjected to a harsh environment of radiation and high magnetic field.  All will be inaccessible once ECAL is assembled.  Approximately half will operate at -10(C; the rest will be maintained at +18(C.

A diagram of the short fanout and long ribbon showing dimensions is shown in Figure 2.

Figure 2.  Diagram of the short fanout and long ribbon.

# .  Possible Construction Sequences

There are two possible sequences for the manufacturer to construct the short fanout and long ribbon, described below.  In the case of Sequence A, it is possible that one manufacturer could produce the short fanout while another produces the long ribbon.

## Sequence A: Construction from free-issue harnesses and ribbon

In Sequence A, each short fanout is to be constructed from half of a free-issue “harness” [].  The harness is assembled by Sumitomo from Ericsson ribbon [].  The long ribbon is to be constructed from the same type of free-issue Ericsson ribbon [2].  This sequence is illustrated in Figure 3.

In the case of Sequence A, this document describes the production of the intermediate fanout termination (one termination), both long ribbon terminations (two terminations), and the intermediate adaptor.

## Sequence B: Construction from free-issue long fanouts

In Sequence B, a short fanout and long ribbon are to be constructed from a free-issue “long fanout”.  (A long fanout would be identical to the standard fanout described in Reference [3] except that the non-naked ribbon section would be approximately 25 meters long.)  This sequence is illustrated in Figure 4.

In the case of Sequence B, this document describes the production of the intermediate fanout termination and the intermediate ribbon termination (two terminations), and the intermediate adaptor.

## Intermediate connectors and adaptor

Significant specifications yet to be defined are the choices of intermediate connector types and intermediate adaptor type.  These will be either MPO or MFS.  They will be chosen by Minnesota/CERN.

Figure 3.  Diagram illustrating Sequence A of production of the short fanout and long ribbon.

Figure 4. Diagram illustrating alternate Sequence B of production of the short fanout and long ribbon.

# .  Short Fanout Specifications

The following specifications must be satisfied by all 12 channels of the short fanout.

All specifications are for assemblies at 25(C unless otherwise noted.

Specifications in italics apply to the free-issue fanout or harness assemblies, described in Section 2 and References [1,].  Many of these specifications are drawn directly from those documents.  The manufacturer is not required to test specifications in italics in this procurement.

## Operational Specifications

| # | Operational Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 9.1.1 | Number of channels | 12 |  |  |
| 9.1.2 | Operation wavelength | 1310 | nm |  |
| 9.2.1 | Connector type on MU-SR adaptor end | MU Simplified Plug (sMU) |  | IEC 61754-6 Insertion orientation shall be marked |
| 9.2.2 | Number of mating cycles | 20 |  |  |  | sMU connector |
| 9.2.3 | Random mate insertion loss1 |  |  | 0.6 | dB | sMU connector |
| 9.2.4 | Random mate return loss1 | 45 |  |  | dB | sMU connector |
| 9.3.1 | Intermediate fanout connector type | MPO or MFS (TBD) |  | To be determined by Minnesota/CERN. |
| 9.3.2 | Number of mating cycles | 20 |  |  |  | Intermediate fanout connector |
| 9.3.3 | Random mate insertion loss1 |  |  | 1.2 | dB | Intermediate fanout connector |
| 9.3.4 | Random mate return loss1 | 50 |  |  | dB | Intermediate fanout connector |
| 9.5.1 | Tensile load, end-to-end |  |  | 10 | N | Operation and installation |
| 9.5.2 | Rotation (on ribbon axis) of intermediate fanout connector with respect to MU-SR adapter |  |  | 0.5 | Turns | Operation and installation |

1Insertion loss and return loss are defined in Reference [3].

## Geometric Specifications

| # | Geometric Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 9.4.1 | Overall fanout length Lfanout |  |  |  |  | Not used: Lfanout is constrained by specifying Lfs and Lfr. |
| 9.4.2 | Single-fiber section length Lfs | 30 | 50 | 100 | cm | Lfs to be distributed in standard lengths with tolerances +3cm/ -0cm. Protected by 900μm loose tube. Set of standard lengths: 30cm, 40cm, 50cm, 60cm, 70cm, 80cm, 90cm, 100cm TBD. |
| 9.4.3 | Ribbon section length Lfr | 128 | 130 | 133 | mm | Approximately 30% of short fanouts (“Preshower type”) have length 130mm +3mm/-2mm. |
|  |  | 1 | 2 | 4 | m | Remaining 70% of short fanouts (“Endcap type”) are to be assembled with Lfr distributed in standard lengths with tolerances +5cm/-0cm. Set of standard lengths: 1.0m, 1.15m, 1.3m, …in 0.15m steps until …3.85m, 4.0m TBD. |
| 9.4.4 | Naked ribbon section length Lfn | 0 | mm | With sheath extending into connector shell1 |
|  |  | 100 |  | Lfr | mm | Alternate possibility at the discretion of the manufacturer. |
| 9.4.5 | Furcation element size | 25 x 10 x 5 | mm |  |

1 Sheath extending just up to the connector is not acceptable.  The sheath and connector boot or shell must provide relief from sharp localized ribbon bending, or else the naked ribbon length must satisfy 100mm < Lfn < Lfr.

## Environmental Specifications

| # | Environmental Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 9.6.1 | Magnetic field |  |  | 4 | T | Parallel to any axis |
| 9.6.2 | Hadronic fluence1 |  |  | 1x1014 | cm-2 | Integrated over lifetime2 |
| 9.6.3 | Gamma radiation dose1 |  |  | 1.5x105 | Gy (Si) | Integrated over lifetime2 |
| 9.6.4 | Temperature | -10 | -10 or 18 | 70 | °C | Operation and storage |
| 9.6.5 | Operating humidity |  | TBD |  |  |  |

1Component resistance to radiation will be verified under the sole responsibility of Minnesota and CERN.
2Foreseen operating lifetime:  nominal 10 years.

## Safety Specifications

| # | Safety Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 9.6.6 | Material composition | Halogen-free, flame retardant material |  | CERN IS41, see reference [] |

## Other Characteristics

### Intermediate Fanout Connector

Connector color	TBD

Ribbon orientation	Index on connector side shall indicate fiber no. 1.  Numbering according to Bellcore color code (fiber 1: dark blue).

Connector boot	To be determined by Minnesota/CERN.

Dust protection	Connector to be protected with dust cap.

Labelling	Free-issue QR-code die to be fixed on connector body.

### Test documentation and traceability

Test file to be delivered at least in electronic form.  File to include at least serial number (if different from QR-code), QR-code numbers (2 values, one for each assembly end), fiber number (1-12), insertion loss (12 values), nominal lengths Lfr (1 value) and Lfn (1 value), label showing nominal length Lfr (1 value, in meters) at intermediate fanout connector end of the assembly.

### Shipping and storage requirements

Shipping	Packaged in cardboard box with appropriate protection of connectors, single-fiber and naked ribbon sections.

Storage	According to best practice.

## Testing

| # | Specification to be tested | Manufacturer | Minnesota/CERN |
| --- | --- | --- | --- |
|  |  | Product Qualification | Lot validation (before delivery) | Pre-production Qualification | Lot Acceptance |
| 9.1.1 | Number of channels | ♦ | ♦ | ♦ | ♦ |
| 9.1.2 | Operation wavelength1 | ♦ | ♦ | ♦ | ♦ |
| 9.2.1 | Connector type on MU-SR adaptor end |  |  | ♦ | ♦ |
| 9.2.2 | Number of mating cycles (sMU) |  |  |  |  |
| 9.2.3 | Random mate insertion loss2 (sMU) |  |  | ♦ | ♦ |
| 9.2.4 | Random mate return loss2 (sMU) |  |  | ♦ | ♦ |
| 9.3.1 | Intermediate fanout connector type | ♦ | ♦ | ♦ | ♦ |
| 9.3.2 | Number of mating cycles (interm. conn.) | ♦ |  | ♦ |  |
| 9.3.3 | Random mate insertion loss2 (interm.) | ♦ | ♦ | ♦ | ♦ |
| 9.3.4 | Random mate return loss2 (interm.) | ♦ | ♦ | ♦ | ♦ |
| 9.5.1 | Tensile load, end-to-end | ♦ |  | ♦ |  |
| 9.5.2 | Rotation (on ribbon axis) of intermediate fanout connector with respect to MU-SR adapter | ♦ |  | ♦ |  |
| 9.4.2 | Single-fiber section length Lfs |  |  | ♦ | ♦ |
| 9.4.3 | Ribbon section length Lfr | ♦ | ♦ | ♦ | ♦ |
| 9.4.4 | Naked ribbon section length Lfn | ♦ | ♦ | ♦ | ♦ |
| 9.4.5 | Furcation element size |  |  | ♦ | ♦ |
| 9.6.1 | Magnetic field |  |  | ♦ |  |
| 9.6.2 | Hadronic fluence |  |  | ♦ |  |
| 9.6.3 | Gamma radiation dose |  |  | ♦ |  |
| 9.6.4 | Temperature | ♦ |  |  |  |
| 9.6.5 | Operating humidity | ♦ |  |  |  |
| 9.6.6 | Material composition | ♦ | ♦ |  |  |

1Wavelength to be used for testing is ~1310 nm.
2Insertion loss and return loss are defined in Reference [3].

## Implementation

The final manufacturer engineering drawing(s) is/are to be inserted here.

# .  Long Ribbon Specifications

The following specifications must be satisfied by all 12 channels of the long ribbon.

All specifications are for assemblies at 25(C unless otherwise noted.

Specifications in italics apply to the free-issue fanout assemblies, described in Section 2 and Reference [3].  Many of these specifications are drawn directly from that document.  The manufacturer is not required to test specifications in italics in this procurement.

## Operational Specifications

| # | Operational Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 10.1.1 | Number of channels | 12 |  |  |
| 10.1.2 | Operation wavelength | 1310 | nm |  |
| 10.2.1 | Intermediate ribbon connector type | MPO or MFS (TBD) |  | To be determined by Minnesota/CERN. |
| 10.2.2 | Number of mating cycles | 20 |  |  |  | Intermediate ribbon connector |
| 10.2.3 | Random mate insertion loss1 |  |  | 1.2 | dB | Intermediate ribbon connector |
| 10.2.4 | Random mate return loss1 | 45 |  |  | dB | Intermediate ribbon connector |
| 10.3.1 | Connector type on end opposite the intermediate ribbon connector | MF-A2 |  | Diamond proprietary2 |
| 10.3.2 | Number of mating cycles | 20 |  |  |  | MF-A2 connector2 |
| 10.3.3 | Random mate insertion loss1 |  |  | 1.2 | dB | MF-A2 connector2 |
| 10.3.4 | Random mate return loss1 | 50 |  |  | dB | MF-A2 connector2 |
| 10.5.1 | Tensile load, end-to-end |  |  | 10 | N | Operation and installation |
| 10.5.2 | Rotation (on ribbon axis) of one connector with respect to other, near connector terminations |  |  | 1 | Turns/m | Operation and installation |

1Insertion loss and return loss are defined in Reference [3].
2In the case of Sequence B (described in Section 2.2), the manufacturer does not need to test specifications 10.3.1 through 10.3.4 in this procurement.

## Geometric Specifications

| # | Geometric Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 10.4.1 | Overall ribbon length Lribbon | 15 | 25 | 35 | m | Lribbon to be distributed in standard lengths with tolerances +10cm/ -0cm. Set of standard lengths: 15.0m, 15.2m, 15.4m, …, 34.8m, 35.0m TBD. |
| 10.4.4 | Naked ribbon section length Lrn | 0 | mm | With sheath extending into connector shell1 |
|  |  | 100 | 100 | 300 | mm | Alternate possibility at the discretion of the manufacturer. |

1Sheath extending just up to the connector is not acceptable.  The sheath and connector boot or shell must provide relief from sharp localized ribbon bending, or else the naked ribbon length must satisfy 100mm < Lrn < 300mm.

## Environmental Specifications

| # | Environmental Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 10.6.1 | Magnetic field |  |  | 4 | T | Parallel to any axis |
| 10.6.2 | Hadronic fluence1 |  |  | 1x1014 | cm-2 | Integrated over lifetime2 |
| 10.6.3 | Gamma radiation dose1 |  |  | 1.5x105 | Gy (Si) | Integrated over lifetime2 |
| 10.6.4 | Temperature | -10 | -10 or 18 | 70 | °C | Operation and storage |
| 10.6.5 | Operating humidity |  | TBD |  |  |  |

1Component resistance to radiation will be verified under the sole responsibility of Minnesota and CERN.
2Foreseen operating lifetime:  nominal 10 years.

## Safety Specifications

| # | Safety Specification | Min | Typ | Max | Units | Note |
| --- | --- | --- | --- | --- | --- | --- |
| 10.6.6 | Material composition | Halogen-free, flame retardant material |  | CERN IS41, see reference [4] |

## Other Characteristics

### Intermediate Ribbon Connector

Connector color	TBD

Ribbon orientation	Index on connector side shall indicate fiber no. 1.  Numbering according to Bellcore color code (fiber 1: dark blue).

Connector boot	To be determined by Minnesota/CERN.

Dust protection	Connector to be protected with dust cap.

Labelling	Free-issue QR-code die to be fixed on connector body.

### Test documentation and traceability

Test file to be delivered at least in electronic form.  File to include at least serial number (if different from QR-code), QR-code numbers (2 values, one for each assembly end), fiber number (1-12), insertion loss (12 values), nominal lengths Lribbon (1 value) and Lrn (1 value), label showing nominal length Lribbon (1 value, in meters) at intermediate ribbon connector end of the assembly.

### Shipping and storage requirements

Shipping	Packaged in cardboard box with appropriate protection of connectors and naked ribbon sections.

Storage	According to best practice.

## Testing

| # | Specification to be tested | Manufacturer | Minnesota/CERN |
| --- | --- | --- | --- |
|  |  | Product Qualification | Lot validation (before delivery) | Pre-production Qualification | Lot Acceptance |
| 10.1.1 | Number of channels | ♦ | ♦ | ♦ | ♦ |
| 10.1.2 | Operation wavelength1 | ♦ | ♦ | ♦ | ♦ |
| 10.2.1 | Intermediate ribbon connector type | ♦ | ♦ | ♦ | ♦ |
| 10.2.2 | Number of mating cycles (intermediate ribbon connector) | ♦ |  | ♦ |  |
| 10.2.3 | Random mate insertion loss2 (intermediate ribbon connector) | ♦ | ♦ | ♦ | ♦ |
| 10.2.4 | Random mate return loss2 (intermediate ribbon connector) | ♦ | ♦ | ♦ | ♦ |
| 10.3.1 | Connector type on end opposite the intermediate ribbon connector (MF-A2)3 | ♦ | ♦ | ♦ | ♦ |
| 10.3.2 | Number of mating cycles (MF-A2)3 | ♦ |  | ♦ |  |
| 10.3.3 | Random mate insertion loss2 (MF-A2)3 | ♦ | ♦ | ♦ | ♦ |
| 10.3.4 | Random mate return loss2 (MF-A2)3 | ♦ | ♦ | ♦ | ♦ |
| 10.5.1 | Tensile load, end-to-end | ♦ |  | ♦ |  |
| 10.5.2 | Rotation (on ribbon axis) of one connector with respect to other, near connector terminations | ♦ |  | ♦ |  |
| 10.4.1 | Overall ribbon length Lribbon | ♦ | ♦ | ♦ | ♦ |
| 10.4.4 | Naked ribbon section length Lrn | ♦ | ♦ | ♦ | ♦ |
| 10.6.1 | Magnetic field |  |  | ♦ |  |
| 10.6.2 | Hadronic fluence |  |  | ♦ |  |
| 10.6.3 | Gamma radiation dose |  |  | ♦ |  |
| 10.6.4 | Temperature | ♦ |  |  |  |
| 10.6.5 | Operating humidity | ♦ |  |  |  |
| 10.6.6 | Material composition | ♦ | ♦ |  |  |

1Wavelength to be used for testing is ~1310 nm.
2Insertion loss and return loss are defined in Reference [3].
3In the case of Sequence B (section ), the manufacturer does not need to test specifications 10.3.1 through 10.3.4.

## Implementation

The final manufacturer engineering drawing(s) is/are to be inserted here.

# .  References

# Endnotes

[]  F. Vasey, “CMS Tracker Optical Readout Link Specification, Part 3.1: Terminated Ruggedized Ribbon Harness”, v. 3.5, CERN EP/CME (2002).  http://edms.cern.ch/document/337825/3.5

[]  F. Vasey, “CMS Tracker Optical Readout Link Specification, Part 3.1.1: Ruggedized Ribbon”, v. 3.7, CERN EP/CME (2002). http://edms.cern.ch/document/312742/3.7

[]  F. Vasey, “CMS Tracker Optical Readout Link Specification, Part 3: Terminated Fiber Ribbon”, v. 3.5, CERN EP/CME (2002).  http://edms.cern.ch/file/344177/3.5/TerminatedRibbonV3_5.pdf

[]  CERN TIS Commission, “The Use of Plastic and other Non-Metallic Materials at CERN with respect to Fire Safety and Radiation Resistance”, CERN TIS IS 41 (1995). http://edms.cern.ch/file/335806/1.01/IS41_E.pdf
http://edms.cern.ch/document/108727/2

# Text boxes

Ribbon
section Lfr

sMU connector

1

1

12

12

1

12

Lribbon

Lfanout

Naked
ribbon 
section Lrn

1

12

12

Naked
ribbon 
section Lfn

Furcation
element

Single fiber
section Lfs

Intermediate
adaptor

MF-A2
connector

1

1

12

Short fanout and long ribbon production (this procurement)

Intermediate ribbon connector

Intermediate fanout connector

Long ribbon

MU-SR
adaptor

Short fanout

12

12

Digital amp. ASIC

PIN photo-diode array

Off Detector

12

Rx module

Intermediate
adaptor

MFS
adaptor

1

Long
ribbon

12

Intermediate Patch Panel

MU-SR
adaptor

CMS

In-Line Patch Panel

Laser diode

Laser driver ASIC

Front End

1

Opto-hybrid Tx

Dense multi-ribbon cable

Short
fanout

Pigtail fiber

Back-end Patch Panel

Distributed Patch Panel

96

12

12

Long ribbon production 
(this procurement)

12

Short fanout production 
(this procurement)

Harness production (Sumitomo)

Fiber ribbon (Ericsson)

12

1

1

12

Long fanout production (Diamond)

Harness production (Sumitomo)

Fiber ribbon (Ericsson)

12