Terminated Multi-Ribbon cable

Lot acceptance Report

e3

26 January 2005

# Overview

In this document we report on the lot acceptance tests carried out on the 5 sampled multi-ribbons. The samples were delivered as batch e3. Table 1 shows the sample history and Table 2 the list of the delivery cable number.

| Batch delivery | Delivery Date | Quantity Delivered | Length | Quantity Sampled | Diamond Cable ID 30201210000XXX |
| --- | --- | --- | --- | --- | --- |
| e3 | 31.Dec.2004 | 10 | 45 m | 10 | 151, 152, 153, 154. 156 157, 158, 159, 160, 170 |

Table 1: sample history

| Cable number: | 30201210000XXX |
| --- | --- |
| 1 | 159 |
| 2 | 154 |
| 3 | 156 |
| 4 | 158 |
| 5 | 160 |
| 6 | 170 |
| 7 | 153 |
| 8 | 151 |
| 9 | 157 |
| 10 | 152 |

Table 2: cable number/ID

# Production tests and results

The procedure for production qualification is described in document CMS-TK-QP-0012. The sequence followed for this acceptance test is illustrated in Figure 1. The specifications tested are shown in Table 3.

The 8 selected terminated fibre ribbons underwent visual inspection, geometrical tests and optical tests

.

Fig.1:  Flow chart of the procedure.

| # | Specification to be tested | Test Specifications | CERN Testing |
| --- | --- | --- | --- |
|  |  | Min | Typ | Max | Units | Production Acceptance |
| 3.1 | Number of channels | 96 |  | ( |
| 3.2 | Operation Wavelength | 1310 | nm |  |
| 3.3 | Connector type (In-line) | MFS12-B1 |  | ( |
|  | MFS12 guide pins | Non-removable, ceramic |  |  |
| 3.4 | Number of mating cycles | 20 |  |  |  |  |
| 3.5 | Random mate insertion loss |  |  | 1.2 | dB | ( |
| 3.6 | Return loss | 50 |  |  | dB | ( |
|  | Connector type (Back-end) | MPO 12 |  | ( |
|  | MPO12 guide pins | Non-removable, standard |  | ( |
| 3.4 | Number of mating cycles | 20 |  |  |  |  |
| 3.5 | Random mate insertion loss |  |  | 1.2 |  | ( |
| 3.6 | Return loss | 50 |  |  |  | ( |
| 3.7 | Tensile load on connector |  |  | 10 | N |  |
| 3.9 | In-line ribbon section length LRinline | 600 (+40/-0) | mm | ( |
|  | Multi-ribbon cable length LCable | 45 (+0.5/-0) | m |  |
|  | Multi-ribbon cable length LRBackend | As per drawing (+-15) | mm | ( |

Legend:

( Visual Inspection	( Geometrical measurement	( Optical Test

Table 3:  Validation Programme overview.

## Visual inspection

- All items are properly packaged.

- The QR labels were properly attached according to the specification and could be read without trouble.

- Cable marking: the figure 1 shows the specifications for each connector-type.

Fig. 2: cable marking

- All cables are the same type:  0407 GAY GQBD

- All QR codes ID are as specified.

- Ribbon 1 always longest on MPO side

- Fibre 1 always on the same side for all ribbons and all cables.

The rest of visual inspection is described below for each cable. The summary is shown in Table 4.

Naked ribbon length measurement:

Length for MPO ribbon has been measured as illustrated on figure 3, included the 3.70cm of the furcation element.

Fig. 3: ribbon length measurement

Cable ID 30201220000151

- Cable mreading_mfs:  1673	Diamond:  1674

- Cable mreading_mpo: 1630	Diamond:  1630

- Naked ribbon length:

|  | MFS | MPO |
| --- | --- | --- |
| Ribbon 1 | 65.0 cm | 85.5 cm |
| Ribbon 8 | 61.5 cm | 55.0 cm |

- Cable marking and ribbons orientation:

MPO		    MFS

Cable ID 30201220000152

- Cable mreading_mfs:  1679	Diamond:  1679

- Cable mreading_mpo: 1723	Diamond:  1723

( m-marker: inverted MFS/MPO

- Naked ribbon length:

|  | MFS | MPO |
| --- | --- | --- |
| Ribbon 1 | 62.5 cm | 88.5 cm |
| Ribbon 8 | 61.0 cm | 55.0 cm |

- Cable marking and ribbons orientation:

MPO		       	  MFS

Cable ID 30201220000153

- Cable mreading_mfs:  2030	Diamond:  2031

- Cable mreading_mpo: 1987	Diamond:  1987

- Naked ribbon length:

|  | MFS | MPO |
| --- | --- | --- |
| Ribbon 1 | 63.0 cm | 87.0 cm |
| Ribbon 8 | 62.0 cm | 55.0 cm |

- Cable marking and ribbons orientation:

MPO		       	        MFS

Cable ID 30201220000154

- Cable mreading_mfs:  2132	Diamond:  2132

- Cable mreading_mpo: 2089	Diamond:  2089

- Naked ribbon length:

|  | MFS | MPO | MPO length Diamond |
| --- | --- | --- | --- |
| Ribbon 1 | 62.0 cm | 85.5 cm |  |
| Ribbon 8 | 63.0 cm | 56.0 cm | 53.2 cm |

- Cable marking and ribbons orientation:

MPO		                      MFS

Cable ID 30201220000156

- Cable mreading_mfs:  1834	Diamond:  1834

- Cable mreading_mpo: 1876	Diamond:  1876

( m-marker: inverted MFS/MPO

- Naked ribbon length:

|  | MFS | MPO | MPO length Diamond |
| --- | --- | --- | --- |
| Ribbon 1 | 64 cm | 82.5 cm | 83.5 cm |
| Ribbon 8 | 61 cm | 52 cm | 51.5 cm |

- Cable marking and ribbons orientation:

MPO		             MFS

Comments: ribbon n04: ferrule broken. See Figures below:

Cable ID 30201220000157

- Cable mreading_mfs:  1934	Diamond:  1934

- Cable mreading_mpo: 1978	Diamond:  1978

( m-marker: inverted MFS/MPO

- Naked ribbon length:

|  | MFS | MPO |
| --- | --- | --- |
| Ribbon 1 | 64.5 cm | 86.0 cm |
| Ribbon 8 | 64.0 cm | 55.0 cm |

- Cable marking and ribbons orientation:

MPO		        MFS

Cable ID 30201220000158

- Cable mreading_mfs:  1883	Diamond:  1883

- Cable mreading_mpo: 1927	Diamond:  1927

( m-marker: inverted MFS/MPO

- Naked ribbon length:

|  | MFS | MPO | MPO length Diamond |
| --- | --- | --- | --- |
| Ribbon 1 | 65.0 cm | 86.5 cm | 83.2 cm |
| Ribbon 8 | 63.5 cm | 56.5 cm | 51.5 cm |

- Cable marking and ribbons orientation:

MPO		       	     MFS

Cable ID 30201220000159

- Cable mreading_mfs:  2081	Diamond:  2080

- Cable mreading_mpo: 2038	Diamond:  2038

- Naked ribbon length:

|  | MFS | MPO | MPO length Diamond |
| --- | --- | --- | --- |
| Ribbon 1 | 63.5 cm | 86.0 cm | 82.9 cm |
| Ribbon 8 | 62.5 cm | 55.5 cm |  |

- Cable marking and ribbons orientation:

MPO	        MFS

Cable ID 30201220000160

- Cable mreading_mfs:  1782	Diamond:  2080

- Cable mreading_mpo: 1825	Diamond:  2038

( m-marker: inverted MFS/MPO

- Naked ribbon length:

|  | MFS | MPO | MPO length Diamond |
| --- | --- | --- | --- |
| Ribbon 1 | 113 cm | 85.5 cm | 83.8cm |
| Ribbon 8 | 112 cm | 55.0 cm |  |

MFS length: 111.1 cm in order to arrange the 900 turned fiber

- Cable marking and ribbons orientation:

MPO	   MFS

Comments: ribbons no 1 and 2 were twisted of 3600 on adapter side.

Cable ID 30201220000170

- Cable mreading_mfs:  1775	Diamond:  1776

- Cable mreading_mpo: 1732	Diamond:  1732

- Naked ribbon length:

|  | MFS | MPO |
| --- | --- | --- |
| Ribbon 1 | 63.0 cm | 116.5 cm |
| Ribbon 8 | 61.5 cm | 86.0 cm |

- Cable marking and ribbons orientation:

MPO		       	       MFS

Summary:

| Cable ID | N0 Cable | (mreading_mfs) -(mreading_mpo) 1 | Stack orientation | Nacked Ribbon length (cm) |
| --- | --- | --- | --- | --- |
|  |  |  | MPO | MFS | MPO ribbon1 | MPO ribbon8 | MFS ribbon1 | MFS ribbon1 |
| 151 | 8 | 43 | 1800 | 1800 | 85.5 | 55 | 65 | 61.5 |
| 152 | 10 | - 44 | 900 | 1800 | 88.5 | 55 | 62.5 | 61 |
| 153 | 7 | 43 | 1800 | 1800 | 87 | 55 | 63 | 62 |
| 154 | 2 | 43 | √ | √ | 85.5 | 56 | 62 | 63 |
| 156 | 3 | - 42 | 900 | 1800 | 82.5 | 52 | 64 | 61 |
| 157 | 9 | - 44 | 1800 | √ | 86 | 55 | 64.5 | 64 |
| 158 | 4 | - 44 | 1800 | √ | 86.5 | 56.5 | 65 | 63.5 |
| 159 | 1 | 43 | 450 | 1800 | 86 | 55.5 | 63.5 | 62.5 |
| 160 | 5 | - 43 | 450 | 1800 | 85.5 | 55 | 113 | 112 |
| 170 | 6 | 43 | √ | √ | 116.5 | 86 | 63 | 61.5 |

Table 4: visual inspection summary

1: - mean inverted MFS/MPO length marker.

## Optical measurements

Optical properties were measured using the setup shown in Fig.4 and 5 for respectively MFS connectors’ measurement and MPO.  For measurement, single MFS connection in adaptor has been used.

Fig.4: Set up for optical measurement of MFS connectors.

Fig.5: Set up for optical measurement of MPO connectors.

2.3.1.	Random mate insertion loss (all fibre channels).

The insertion loss (IL) is defined as the Log of the ratio of optical powers measured before (P0) and after (P1) insertion of the device under test.  The optical powers (P0 and P1) were measured with a large area detector in an optical head receptacle.

Insertion loss:

#### 2.3.1.a	Acceptance criteria

The acceptance criterion is such that the insertion loss of the MFS connector shall be less than 1.2 dB.

#### 2.3.1.b	Results

The results are shown in Fig. 6 for MFS side and Fig.7 for MPO side.

Fig.6: Insertion loss of MFS connectors.

Fig.7 Insertion loss of MPO connectors

On cable 160, 3 channels from ribbon 1 are higher than 1.2 db else all others cables are in specifications.

For these measurements, some values were higher than 1.2 dB time to time but after cleaning connectors, values were fine.

Nevertheless for some ribbons, it was very difficult to get IL<1.2db on MFS side only: for 4 ribbons of different cables, almost 10 re-mates with cleaning have been necessary to be in specifications.

Figure on the left shows you an example of several re-mates on ribbon 1 of cable 160. Despite cleaning, there was no improvement for this ribbon.

On figure 8 and 9 is shown the histogram for all the MFS and MPO insertion loss values.

Fig.8:  Histogram of IL values for MFS connectors

Fig.9:  Histogram of IL values for MPO connectors

### 2.3.2.	Return loss (all fibre channels).

The return loss (RL) is defined as the Log of the ratio of the reflected optical powers measured during referencing (Pref) and after (P1) insertion of the device under test (DUT).

Return loss: (1)

#### 2.3.2.a	Acceptance criteria

The acceptance criterion is such that the return loss shall be greater than 50dB.

#### 2.3.2.b	Results

The results are shown in Fig. 10 and 11; respectively for MFS and MPO sides. All devices were well within the specified limit.

Fig.10: Return loss of MFS connectors.

Fig.11: Return loss of MPO connectors.

# Conclusions

# Footnotes

(1) Formula valid if using a reference with perfect glass-air interface.

# Text boxes

( Ribbon stack inverted 	(=1800)

( Ribbon stack inverted

Ribbon 1

( BAD orientation, 450

(Ribbon stack inverted

Ribbon 1

Cable-marking

( Ribbon stack inverted

( BAD orientation (900)

Reception

Ribbon 1

Cable-marking

Cable-marking

Cable-marking

( BAD orientation (900)

Ribbon 1

Ribbon 1

( Ribbon stack inverted

Cable-marking

( Ribbon stack inverted

( Ribbon stack inverted

Cable-marking

Ribbon 1

( OK

( OK

Cable-marking

Cable-marking

Ribbon 1

Ribbon 1

( Ribbon stack inverted

( OK

Cable 160, ribbon1

Ribbon 1

Ribbon 1

(Ribbon stack inverted

(OK

Cable-marking

Ribbon 1

(Ribbon stack inverted

( BAD orientation, 450

Ribbon 1

Cable-marking

Ribbon 1 MPO

(OK

(OK